## Summary
I_designed [A] | U_designed [%] | I_measured [A] | U_measured [%]
0.5 | 0.13641
0.05 | 0.079052 | 0.049998 | 0.0041388
0.00125 | 0.078947 | 0.0012498 | 0.014864
0.0005 | 0.075763 | 0.00049998 | 0.0040509
5e-05 | 0.074994 | 5e-05 | 0.00097889
1.25e-06 | 0.11463 | 1.2507e-06 | 0.05317
5e-07 | 0.11463 | 5.007e-07 | 0.13961
5e-08 | 0.16473 | 5.0639e-08 | 1.2784
1e-08 | 0.48661 | 1.0648e-08 | 6.4808

## Measurements
I_designed [A] | I_measured [A] | Error [%] | Standard dev [A] | Uncertainty [A] | Uncertainty [%]
0.5
0.05 | 0.049998 | -0.0041388 | 1.4928e-07 | 2.9999e-05 | 0.06
0.00125 | 0.0012498 | -0.014864 | 3.6508e-08 | 2.6249e-06 | 0.21002
0.0005 | 0.00049998 | -0.0040509 | 1.0013e-09 | 2.9999e-07 | 0.06
5e-05 | 5e-05 | -0.00097889 | 1.3231e-10 | 2.6e-08 | 0.052
1.25e-06 | 1.2507e-06 | 0.05317 | 9.7927e-11 | 8.2533e-10 | 0.065991
5e-07 | 5.007e-07 | 0.13961 | 9.602e-11 | 3.0035e-10 | 0.059986
5e-08 | 5.0639e-08 | 1.2784 | 9.5511e-11 | 7.532e-11 | 0.14874
1e-08 | 1.0648e-08 | 6.4808 | 9.2154e-11 | 5.5324e-11 | 0.51957
All measurements were made with a Keysight 34465A with a calibration date of less than 1 year.
The 34465A was warmed up for at least 1 hour before measurements and then an autocalibration was performed.
All measurements were taken with 10 NPLC, auto zero, and the lowest current range possible.
The DUT reading was allowed to stabilize for 2 minutes (typically took less than 30 seconds) before data recording.
500 data points were taken for each current setting (~3 minutes).
The measurement uncertainty was calculated using the 1-year values specified in the 34465A manual.

## Designed tolerance
[6 empty] | Resistors |  |  | Voltage reference | [5 empty] | Trace resistance |  | Op amp
V_ref [V] | R [Ω] | I [A] | U [%] |  | T [C] | Total tol [%] | R_tol [%] | T_coeff [ppm/C] | Total tol [%] | Base tol [%] | T drift [%] | Line reg [μV/V] | Load reg [μV/mA] | Long-term [%] | Sense offset [%] | Gnd offset [%] | Total tol [%] | Vos [%] | T drift [%] | Ib [%] | Ios [%]
1.25 | 2.5 | 0.5 | 0.13641 | 0.50068 | 50 | 0.055902 | 0.05 | 10 | 0.066283 | 0.04 | 0.015 | 0.05 | 0.0075 | 0.0035 | 0.0318 | 0.027695 | 0.00065106 | 0.00064 | 0.0001 | 1.36e-05 | 6.4e-05
| 25 | 0.05 | 0.079052 | 0.05004 | 35 | 0.01118 | 0.01 | 5 | 0.064842 | 0.04 | 0.006 | 0.05 | 0.0075 | 0.0035 | 0.00378 | 0.0078877 | 0.00064458 | 0.00064 | 4e-05 | 1.36e-05 | 6.4e-05
| 1000 | 0.00125 | 0.078947 | 0.001251 | 25 | 0.01 | 0.01 | 5 | 0.064564 | 0.04 | 0 | 0.05 | 0.0075 | 0.0035 | 0.0103 | 0.00059468 | 0.00064334 | 0.00064 | 0 | 1.36e-05 | 6.4e-05
| 2500 | 0.0005 | 0.075763 | 0.00050038 | 25 | 0.010396 | 0.010396 | 5.078 | 0.064564 | 0.04 | 0 | 0.05 | 0.0075 | 0.0035 | 0.004104 | 0.00048248 | 0.00064334 | 0.00064 | 0 | 1.36e-05 | 6.4e-05
| 25000 | 5e-05 | 0.074994 | 5.0037e-05 | 25 | 0.010396 | 0.010396 | 5.078 | 0.064564 | 0.04 | 0 | 0.05 | 0.0075 | 0.0035 | 0.0004304 | 0.00055088 | 0.00064334 | 0.00064 | 0 | 1.36e-05 | 6.4e-05
| 1000000 | 1.25e-06 | 0.11463 | 1.2514e-06 | 25 | 0.05 | 0.05 | 10 | 0.064564 | 0.04 | 0 | 0.05 | 0.0075 | 0.0035 | 1.089e-05 | 0.00044078 | 0.00064334 | 0.00064 | 0 | 1.36e-05 | 6.4e-05
| 2500000 | 5e-07 | 0.11463 | 5.0057e-07 | 25 | 0.05 | 0.05 | 10 | 0.064564 | 0.04 | 0 | 0.05 | 0.0075 | 0.0035 | 4.296e-06 | 0.00040775 | 0.00064334 | 0.00064 | 0 | 1.36e-05 | 6.4e-05
| 25000000 | 5e-08 | 0.16473 | 5.0082e-08 | 25 | 0.1 | 0.1 | 25 | 0.064564 | 0.04 | 0 | 0.05 | 0.0075 | 0.0035 | 4.196e-07 | 0.00042224 | 0.00064334 | 0.00064 | 0 | 1.36e-05 | 6.4e-05
| 125000000 | 1e-08 | 0.48661 | 1.0049e-08 | 25 | 0.42 | 0.42 | 50 | 0.064564 | 0.04 | 0 | 0.05 | 0.0075 | 0.0035 | 8.392e-08 | 0.00042224 | 0.00064334 | 0.00064 | 0 | 1.36e-05 | 6.4e-05

## Resistor configuration
V_ref [V] | R [Ω] | I [A] | P [W] | Resistor config | R_tol [%] | P_R [W] | T_coeff [ppm/C] | Package | Part #
1.25 | 2.5 | 0.5 | 0.625 | 4//10R .05% | 0.05 | 1 | 10 | 1206 | RNCF1206AKT10R0
| 25 | 0.05 | 0.0625 | 4//100R .01% | 0.01 | 0.25 | 5 | 0603 | APC0603T100RZ
| 1000 | 0.00125 | 0.0015625 | 1k .01% | 0.01 | 0.1 | 5 | 0805 | APC0805T1K00Z
| 2500 | 0.0005 | 0.000625 | 2k, 499R .01%, 1R 1% | 0.010396 | 0.12 | 5.078 | 1206 0805 0603 | APC1206T2K00Z APC0805T499RZ RC0603FR-071RL
| 25000 | 5e-05 | 6.25e-05 | 2x10k, 4.99k .01%, 10R 1% | 0.010396 | 0.062479 | 5.078 | 0603 0603 0603 | APC0603T10K0Z APC0603T4K99Z RC0603FR-0710RL
| 1000000 | 1.25e-06 | 1.5625e-06 | 1M .05% | 0.05 | 0.25 | 10 | 1206 | RG3216N-1004-W-T1
| 2500000 | 5e-07 | 6.25e-07 | 2x1M, 470k, 30k .05% | 0.05 | 0.2247 | 10 | 1206 0805 0603 | RG3216N-1004-W-T1 RG2012N-474-W-T1 RG1608N-303-W-T1
| 25000000 | 5e-08 | 6.25e-08 | 12x2M, 1M .1% | 0.1 | 1.3 | 25 | 0805 | CPF0805B2M0E1 CPF0805B1M0E1
| 125000000 | 1e-08 | 1.25e-08 | 10x10M .5%, 12x2M, 1M .1% | 0.42 | 1.08 | 50 | 0603 | CRCW060310M0DHEAP
Note on E10: Matthew Blessinger: Cascade 100M into 25M for 125M